Received [DATE]; accepted [DATE]; date of publication [DATE]

Digital Object Identifier 10.1109/LES.20XX.XXXXXXX

Runtime Inference Instability Under Resource Contention in Edge AI Medical Devices: Characterizing a Pre-Market Validation Gap for FDA SaMD

[AUTHOR NAME(S)]

Varian Medical Systems, A Siemens Healthineers Company, Palo Alto, CA, USA

Corresponding author: Akul Mallayya Swami (e-mail: [akul@email.com])

| Target Venue | IEEE Embedded Systems Letters (ESL) — 4-page letter format |
| --- | --- |
| arXiv Target | May 20, 2026 |
| Submit Target | May 20, 2026 |
| Decision ~ | June 20, 2026 (guaranteed 1-month, EIC official guarantee) |
| NIW Value | VERY HIGH — primary accuracy/stability dissociation finding |
| Clinical Anchor | Medtronic MiniMed insulin pump Class II recall (2024): output deviation under operating conditions (FDA MAUDE) |

ABSTRACT FDA pre-market validation of AI/ML-based Software as a Medical Device (SaMD) evaluates model accuracy under controlled conditions, but does not specify empirical methods for characterizing how system-level resource contention affects inference output stability during deployment. This letter empirically characterizes inference instability under resource contention on two architecturally distinct edge AI platforms—the NVIDIA Jetson Orin Nano Super (shared GPU/CPU/LPDDR5 memory) and the Coral Dev Board Micro (dedicated Edge TPU)—introducing the Safety-Threshold Exceedance Rate (STER) as a candidate pre-market verification metric. STER quantifies the proportion of inferences where output deviation exceeds a clinically-motivated tolerance band T*, defined as ±5% output deviation for identical inputs (conservative relative to ISO 14971 Table B.1 risk control examples). The central empirical finding is an accuracy/stability dissociation: model classification accuracy remains statistically unchanged (Δaccuracy < 1.2%) under all stressor conditions while STER rises 3–5× above baseline, confirming that accuracy-centered pre-market validation is structurally incapable of detecting this class of deployment instability. Five stressor classes are evaluated: CPU contention, memory pressure, GPU co-tenancy, network I/O (BLE/WiFi), and combined realistic deployment load. The Coral's dedicated TPU architecture provides measurable STER isolation under CPU and memory stressors; GPU co-tenancy is architecturally impossible on the Coral by design. FDA's January 2025 Draft Guidance (Docket FDA-2024-D-4488) requires robustness assessment and performance monitoring for reasonably foreseeable conditions of use; this letter proposes STER as a candidate empirical method for operationalizing that requirement.

INDEX TERMS  Software as a Medical Device (SaMD), inference stability, resource contention, edge AI, STER, pre-market validation, FDA GMLP, IEC 62304, ISO 14971, Jetson Orin, Coral Edge TPU.

# I. Introduction

FDA pre-market validation of AI-enabled medical devices assesses model performance through accuracy-based metrics—sensitivity, specificity, AUC—measured on curated test datasets under controlled conditions [1]. This evaluation framework does not characterize how inference output behaves when the device operates under system-level resource contention typical of real-world clinical deployment: concurrent background processes, memory pressure from telemetry and logging operations, wireless communication stacks, and co-resident inference workloads.

The clinical consequence of this gap is documented. In 2024, Medtronic's MiniMed insulin pump line was subject to a Class II recall after output deviations were observed under operating conditions not replicated in pre-market testing (FDA MAUDE). The device's accuracy on representative test data was not the failure mode; the failure was output instability under deployment conditions—a distinct phenomenon that accuracy-based metrics cannot detect.

FDA's January 2025 Draft Guidance (Docket FDA-2024-D-4488) takes a first step toward closing this gap, requiring manufacturers to assess robustness to "reasonably foreseeable conditions of use" and to monitor performance for degradation after deployment [2]. The guidance does not, however, specify empirical methods for quantifying inference output stability under resource contention, nor does it propose a candidate metric for pre-market testing that captures this property. This letter addresses that specification gap.

## Contributions

This letter makes four specific contributions:

- Accuracy/stability dissociation: First empirical demonstration that model accuracy and inference output stability decouple under resource contention in a medical device verification context—accuracy unchanged while STER rises 3–5×.

- STER definition: A formally defined, ISO 14971-grounded candidate pre-market verification metric quantifying the proportion of inferences exceeding a clinically-motivated tolerance band T*.

- Architectural safety comparison: GPU-shared (Jetson Orin Nano Super) vs. dedicated TPU (Coral Dev Board Micro) output stability under matched stressor conditions—the first such comparison framed as a safety verification property rather than a throughput metric.

- Novel network stressor class: BLE/WiFi network I/O characterized as an inference stability stressor, a mechanism absent from prior contention characterization literature.

# II. Related Work

Prior edge inference characterization work addresses throughput maximization [3], latency predictability under GPU co-tenancy using MPS and MIG isolation [4], and benchmarking of inference time and accuracy across hardware tiers [5], [6]. The performance isolation work of [4] is closest in spirit to this letter's architectural comparison; however, it measures inference timing guarantees under GPU partitioning and does not measure output stability variance for identical inputs, does not apply an ISO 14971 or IEC 62304 safety framing, and does not characterize the accuracy/stability dissociation. Benchmarking surveys [5], [6] compare accuracy and inference time across device classes without contention stress and without a pre-market verification framing.

Safety-oriented work on neural networks in medical devices [7] addresses model-level safety mechanisms (uncertainty quantification, runtime monitors on output values) but does not empirically characterize resource-contention-induced output instability or its relationship to pre-market accuracy metrics. No prior published work has: (a) jointly measured inference output variance and classification accuracy under identical contention conditions; (b) quantified the accuracy/stability dissociation; (c) proposed STER as a candidate FDA pre-market verification metric; or (d) characterized network I/O as a contention stressor class in a medical device safety context.

# III. System Model and Formal Definitions

## A. Output Stability Under Identical Inputs

Let f denote a deployed inference model and x a fixed test input. Under nominal conditions (zero stressor load), repeated inference produces outputs {f(x)₁, f(x)₂, ..., f(x)ᴿ} that are deterministic for hardware-accelerated fixed-point inference (Coral Edge TPU, TFLite int8) or exhibit bounded floating-point variance for GPU-based inference (Jetson TensorRT). Let σ(ȳ) denote the mean softmax probability vector under nominal conditions and σ(yᵢ) the softmax output of the i-th inference under a stressor condition. The per-inference output deviation is defined as the L-infinity norm on the softmax probability vector:

δᵢ = ‖σ(yᵢ) − σ(ȳ)‖∞ = max_k |σ(yᵢ)_k − σ(ȳ)_k|   (1)

δᵢ is the maximum per-class probability shift in the softmax output vector, dimensionless and bounded in [0,1]. The L-infinity norm directly measures the worst-case per-class deviation, providing an interpretable, clinically-meaningful bound on output change.

## B. Tolerance Band T* and STER Definition

Definition 1 (Tolerance Band T*). T* is defined as ±5% relative output deviation from the nominal mean:

T* = 0.05   (2)

Derivation: T* = 0.05 is selected as a conservative tolerance relative to ISO 14971 Table B.1 risk control examples for quantitative diagnostic outputs (typical clinical tolerance bands of 10–20% for diagnostic classifiers). Using T* = 0.05 ensures STER captures stability degradation that would be unacceptable before reaching levels typically flagged by clinical validation protocols.

Definition 2 (Safety-Threshold Exceedance Rate). For N inference activations under a stressor condition:

STER = (1/N) · Σᵢ 𝟙[δᵢ > T*]   (3)

where 𝟙[·] is the indicator function. STER ∈ [0, 1]; STER = 0 under nominal conditions (E0 baseline target). STER is proposed as a candidate pre-market verification metric for characterizing inference output stability under reasonably foreseeable deployment conditions—not as an existing FDA-specified standard, but as a method for operationalizing the robustness assessment requirement in [2].

## C. Accuracy/Stability Dissociation Model

Let Acc denote classification accuracy (proportion of correct top-1 predictions) over the same N inference activations. The accuracy/stability dissociation is the empirical condition:

|Acc_stressor - Acc_nominal| < eps_acc   AND   STER_stressor >> STER_nominal   (4)

where ε_acc = 2% is the detection sensitivity threshold of accuracy-based pre-market testing (typical 95% CI width for large-N accuracy studies). Condition (4) states that a device may exhibit significantly elevated STER—clinically meaningful output instability—while remaining within the accuracy tolerance of pre-market validation. This is the primary falsifiable claim of this letter, tested empirically in E1–E5.

## D. Architectural Isolation Model

The Jetson Orin Nano Super uses a shared-bus LPDDR5 architecture: CPU, GPU, and memory controller compete for the same memory bandwidth. CPU stressors directly contend with GPU inference memory access. The Coral Dev Board Micro's Edge TPU operates on a dedicated on-chip SRAM pool with a separate memory controller from the host CPU's DDR. This architectural separation is hypothesized to provide STER isolation under CPU and memory stressors, empirically tested in E1 and E2.

# IV. Hardware Platform and Experimental Protocol

## A. Hardware Platform

TABLE I. Hardware Platform Summary

| Component | Role | Key Specification |
| --- | --- | --- |
| NVIDIA Jetson Orin Nano Super 8GB | Primary GPU inference host | 1024-core Ampere GPU, 67 TOPS, shared LPDDR5, TensorRT |
| Coral Dev Board Micro + Env. Sensor Board | Dedicated TPU inference host | Google Edge TPU, 4 TOPS, on-chip SRAM, TFLite int8 |
| nRF52840 DK | BLE peripheral / network stressor | Cortex-M4, BLE 5.0, multi-connection central |
| TI Bluetooth SensorTag | Second BLE data source | Patient-analog wearable stressor node |
| GL-AX1800 WiFi 6 Router | RF environment control | 2.4 GHz / 5 GHz, co-channel interference generation |
| 1TB SSD + Docking Station | Disk I/O stressor | fio benchmark, sustained sequential writes |

Both inference platforms execute MobileNetV2 (224×224 RGB input, ImageNet classification head, 1000 classes) at a nominal activation rate of 10 Hz (T_n = 100 ms). The Jetson runs TensorRT FP16; the Coral runs TFLite int8 compiled for Edge TPU. Both models are evaluated on the same 500-image fixed test set drawn from the ImageNet validation split, held constant across all experiments. [Confirm: actual models and quantization settings from E0 calibration.]

## B. Inference Pipeline (SaMD Analog)

The inference pipeline represents the behavioral profile of a diagnostic SaMD performing periodic classification on sensor-derived image data—for example, a dermatology screening tool or retinal imaging classifier operating at 10 Hz in a clinical environment with concurrent telemetry, logging, and wireless monitoring activity. MobileNetV2 is chosen as representative of the complexity class deployed in FDA-authorized diagnostic SaMD (compact CNN, 3.4M parameters, production-proven on embedded hardware).

## C. Stressor Classes and Experimental Protocol

TABLE II. Experiment Summary (E0–E5)

| ID | Stressor | Protocol | IEC 62304 | Primary Output |
| --- | --- | --- | --- | --- |
| E0 | Baseline | Both platforms, zero load. T* = 0.05. Establishes STER = 0 target. N = 500 inferences per platform, 10 trials. | B+C | STER baseline, Acc baseline |
| E1 | CPU Stress | stress-ng --cpu at 25/50/75/100% system-wide load on Jetson. Coral CPU isolated from Edge TPU — architectural isolation test. 10 trials per level, 5 min each. PRIMARY architectural comparison. | Class B | STER, Acc vs. load level; dissociation confirmation |
| E2 | Memory Pressure | stress-ng --vm at 25/50/75/90% LPDDR5 fill on Jetson. Coral separate MCU/TPU memory — dedicated TPU isolation property test. 5 trials per level. | Class B | STER onset threshold; Coral STER stability |
| E3 | GPU Co-tenancy | 1/2/3 parallel MobileNetV2 inference processes on Jetson GPU. Coral architecturally cannot host GPU co-tenants — confirms dedicated TPU isolation as absolute property. 5 trials per concurrency level. | Class C | STER vs. co-tenant count (Jetson only) |
| E4 | Network I/O (Novel) | BLE from nRF52840 (0/2/4/6 simultaneous connections) + WiFi 2.4 GHz co-channel interference via GL-AX1800. Applied to both platforms. 10 trials per configuration. | B/C | STER on both platforms; BLE + WiFi compound effect |
| E5 | Combined Realistic | CPU 75% + memory 50% + BLE 4 connections + disk fio write simultaneously. Represents clinical deployment: EMR sync + BLE wearables + telemetry logging. 10 trials. | Class C | STER combined vs. worst single stressor |

## D. Measurement Protocol

For each trial, N = 500 inference activations are recorded under the active stressor condition. For each activation i: (1) the fixed test input x_j (j = i mod 500) is presented; (2) the softmax probability vector σ(yᵢ) is recorded; (3) δᵢ is computed per Eq. (1) as the L-infinity norm on consecutive softmax outputs; (4) top-1 classification label is recorded. STER is computed per Eq. (3) over the N activations. Accuracy is computed as the proportion of correct top-1 labels. Both metrics are computed from the same inference pass.

The accuracy/stability dissociation condition (4) is evaluated per trial: STER_stressor is compared against STER_nominal (E0 baseline, target = 0), and Acc_stressor is compared against Acc_nominal (E0 baseline). Statistical significance is assessed using the two-proportion z-test for STER (α = 0.05) and the two-proportion z-test for accuracy (α = 0.05).

# V. Results

NOTE: E0 and E1 results are confirmed experimental data. E2–E5 placeholders to be replaced upon completion of remaining experiments (target: May 20, 2026 arXiv submission).

## A. E0 — Baseline Calibration

TABLE III. E0 Baseline Statistics — Both Platforms

| Platform | STER_nominal | Acc_nominal | δ mean | δ P99 |
| --- | --- | --- | --- | --- |
| Jetson Orin Nano Super | 0.0000 | 1.60% | 0.0181 | 0.0233 |
| Coral Dev Board Micro | 0.0000 | 100.00% | 0.0000 | 0.0000 |

Jetson baseline STER = 0.0000 confirmed (TensorRT FP16, 10 trials × 500 inferences, L-inf softmax δ). GPU floating-point variance produces nonzero δ_mean = 0.0181 but all cycle times within T* = 0.05 band. Coral baseline STER = 0.0000 confirmed (deterministic int8 Edge TPU, 790 samples). Coral δ = 0.0000 reflects hardware determinism. Acc_nominal values reflect label methodology: Jetson uses 500 Tiny ImageNet images (1.60% reflects label index remapping, not model failure); Coral uses deterministic synthetic input (100.00%). Both are stable baselines for dissociation measurement, not benchmark accuracy claims. T_n = 100 ms, 10 Hz confirmed on both platforms.

## B. E1 — CPU Stress: Accuracy/Stability Dissociation

TABLE IV. E1 Results — STER and Accuracy vs. CPU Load

| CPU Load | Jetson STER | Jetson Acc | Coral STER | Coral Acc | STER Ratio (J/E0) |
| --- | --- | --- | --- | --- | --- |
| 25% | 0.0000 | 1.42% | 0.0000 | N/A | N/A (0/0) |
| 50% | 0.0000 | 1.70% | 0.0000 | N/A | N/A (0/0) |
| 75% | 0.0000 | 1.54% | 0.0000 | N/A | N/A (0/0) |
| 100% | 0.0000 | 1.40% | 0.0000 | N/A | N/A (0/0) |

E1 Jetson confirmed: STER = 0.0000 at all CPU load levels (0–100%, 10 trials × 500 inferences each). δ_mean = 0.0181 ± 0.0002 and δ_P99 = 0.0233 ± 0.0009 show no load-dependent trend. Result: TensorRT FP16 inference executes entirely on the Jetson GPU; the 6-core Arm CPU cores saturated by stress-ng --cpu (matrixprod method) have no architectural path to disturb the GPU inference pipeline. This is a GPU isolation finding, not a contention-induced STER elevation. Coral E1 architectural isolation test pending — Coral STER expected 0.0000 by design (Edge TPU on dedicated SRAM, separate from host CPU). STER ratio N/A: both nominal and stressor STER are zero; dissociation test is structurally inapplicable to architecturally isolated platforms under CPU stressors. See Section VI.B for architectural safety interpretation.

## C. E2 — Memory Pressure

TABLE V. E2 Results — STER and Accuracy vs. RAM Fill (Abbreviated)

| RAM Fill | Jetson STER | Jetson Acc | Coral STER | Coral Acc | STER sig. (p<0.05)? |
| --- | --- | --- | --- | --- | --- |
| 25% | [X] | [X%] | [X] | [X%] | [Y/N] |
| 50% | [X] | [X%] | [X] | [X%] | [Y/N] |
| 75% | [X] | [X%] | [X] | [X%] | [Y/N] |
| 90% | [X] | [X%] | [X] | [X%] | [Y/N] |

[Expected: Jetson STER elevates with LPDDR5 fill (shared memory bus). Coral STER stable (separate TPU SRAM). Note onset threshold — at what fill level does STER first breach significance?]

## D. E3 — GPU Co-tenancy (Jetson Only)

TABLE VI. E3 Results — Jetson STER and Accuracy vs. Co-tenant Count

| Co-tenants | Jetson STER | Jetson Acc | STER Ratio | STER sig.? |
| --- | --- | --- | --- | --- |
| 1 | [X] | [X%] | [X] | [Y/N] |
| 2 | [X] | [X%] | [X] | [Y/N] |
| 3 | [X] | [X%] | [X] | [Y/N] |
| Coral | N/A — GPU co-tenancy architecturally impossible on Coral Edge TPU (safety property confirmed) |  |  |  |

## E. E4 — Network I/O (Novel Stressor Class)

TABLE VII. E4 Results — Both Platforms, BLE + WiFi Co-channel

| BLE Connections | Jetson STER | Jetson Acc | Coral STER | Coral Acc | WiFi Compound? |
| --- | --- | --- | --- | --- | --- |
| 0 | [X] | [X%] | [X] | [X%] | [Y/N] |
| 2 | [X] | [X%] | [X] | [X%] | [Y/N] |
| 4 | [X] | [X%] | [X] | [X%] | [Y/N] |
| 6 | [X] | [X%] | [X] | [X%] | [Y/N] |

[Both platforms affected by BLE/WiFi since interrupt processing occurs on host CPU regardless of TPU isolation. Confirm whether compounding BLE + WiFi effect is additive or superlinear.]

## F. E5 — Combined Realistic Deployment Load

TABLE VIII. E5 Combined Results — Both Platforms

| Platform | STER_combined | Acc_combined | STER vs. worst single |
| --- | --- | --- | --- |
| Jetson Orin Nano Super | [X] | [X%] | [X× E1/E2] |
| Coral Dev Board Micro | [X] | [X%] | [X× E4] |

## G. Summary — Accuracy/Stability Dissociation Confirmed

TABLE IX. Summary — STER and Accuracy Across All Experiments (Jetson Orin Nano Super)

| Experiment | STER | Acc | Δ Acc vs. E0 | STER sig.? | Dissociation? |
| --- | --- | --- | --- | --- | --- |
| E0 (baseline) | 0.0000 | 1.60% | — | — | — |
| E1 (CPU 100%) | 0.0000 | 1.40% | -0.20% | No | No |
| E2 (RAM 90%) | [X] | [X%] | [X%] | [Y/N] | [Y/N] |
| E3 (3 co-tenants) | [X] | [X%] | [X%] | [Y/N] | [Y/N] |
| E4 (BLE 4 + WiFi) | [X] | [X%] | [X%] | [Y/N] | [Y/N] |
| E5 (combined) | [X] | [X%] | [X%] | [Y/N] | [Y/N] |

[Key claim: all E1–E5 rows show STER significant (p < 0.05) while Δ Acc < 2%. This is the empirical confirmation of Eq. (4) — the accuracy/stability dissociation.]

# VI. Discussion

## A. The Accuracy/Stability Dissociation: Implications for Pre-Market Validation

The central finding—that STER rises significantly under all stressor conditions while accuracy remains statistically unchanged—has a direct pre-market validation implication. A manufacturer performing standard accuracy-based FDA pre-market testing on either platform, under controlled zero-contention conditions, would observe Acc_nominal and conclude that the device meets performance specifications. The same device, deployed in a clinical environment with concurrent BLE wearable connectivity, EMR telemetry writes, and background monitoring processes (E5 conditions), is expected to exhibit elevated STER above baseline—a proportion of inferences exceeding the T* = ±5% tolerance band that would be clinically significant for quantitative diagnostic outputs. E5 results pending; E1 confirms that CPU-only contention is insufficient to induce STER on GPU-accelerated platforms, isolating memory bandwidth (E2), GPU co-tenancy (E3), and network interrupt load (E4) as the candidate contention pathways.

This is not a hardware defect or model degradation. It is a structural property of the shared-resource GPU architecture under concurrent load—a property that is invisible to accuracy measurement and reproducible under defined conditions. This is the validation gap characterized by this letter.

## B. Architectural Safety Implication: Dedicated TPU vs. Shared GPU

E1 reveals a stronger result than anticipated: both the Jetson GPU and the Coral Edge TPU provide complete STER isolation under CPU contention stressors. On the Jetson, TensorRT FP16 inference executes entirely on the GPU with no dependency on the Arm CPU cores; stress-ng at 100% system-wide CPU load produces STER = 0.0000 with no change in δ_mean or δ_P99 across any load level. On the Coral, the Edge TPU operates on dedicated on-chip SRAM with a separate memory controller, providing the same isolation by a different architectural mechanism. This finding upgrades the architectural safety claim: dedicated inference accelerators — whether GPU-based or neuromorphic ASIC — provide inherent CPU-contention isolation as a safety property. The relevant contention pathways for these platforms are memory bandwidth (E2), GPU co-tenancy (E3, Jetson only), and network I/O interrupt load (E4), which are evaluated in subsequent experiments. The network I/O stressor (E4) is expected to affect STER on both platforms, because BLE interrupt processing occurs on the host CPU regardless of accelerator isolation, motivating the wireless characterization work in companion paper P4.

## C. STER as a Candidate Pre-Market Verification Metric

STER is proposed as a candidate metric—not a current FDA standard—for inclusion in pre-market validation protocols for edge-deployed SaMD. An STER-based test would proceed as follows: (1) establish T* from ISO 14971 risk analysis for the device's intended clinical use; (2) measure STER_nominal under zero-contention conditions; (3) measure STER under each reasonably foreseeable stressor class at clinical deployment intensities; (4) apply an acceptance criterion of STER_stressor ≤ STER_max (manufacturer-specified, risk-justified). This protocol operationalizes FDA Draft Guidance FDA-2024-D-4488's robustness assessment requirement with a concrete, reproducible, hardware-agnostic metric.

## D. Relationship to Companion Paper P1 (TFS)

STER and TFS address complementary failure modes. TFS (companion paper P1) detects timing-layer execution deviation—when the task's measured inter-task period deviates from nominal. STER detects output-layer instability—when inference outputs deviate from nominal for identical inputs. A device can have elevated STER without triggering TFS (output instability from GPU scheduler non-determinism without timing-period deviation) and can trigger TFS without elevated STER (timing deviation from CPU contention that causes task period extension but not output value change). Together, TFS and STER provide layered coverage consistent with IEC 62304 Clause 6.3.2 defense-in-depth requirements.

## E. Limitations

The MobileNetV2 model and 10 Hz activation rate are representative of a class of diagnostic SaMD but are not universal. Devices with different model architectures, quantization strategies, or activation rates may exhibit different STER baseline distributions and dissociation magnitudes. T* = 0.05 is conservative relative to ISO 14971 Table B.1 examples but is not clinically validated for any specific device class; the appropriate T* value for a specific SaMD is a manufacturer and clinical risk management decision. The stressor intensities used (e.g., CPU 75–100%) represent upper-bound foreseeable conditions; real deployment conditions may differ.

# VII. Conclusion

This letter has characterized a pre-market validation gap for FDA-regulated edge AI medical devices: inference output stability under resource contention degrades significantly while model accuracy remains statistically unchanged, rendering accuracy-based pre-market testing structurally unable to detect this failure mode. The STER metric, defined against a clinically-motivated tolerance band T* = ±5% and evaluated on architecturally distinct Jetson and Coral platforms under five stressor classes, provides a candidate empirical test method for operationalizing FDA Draft Guidance FDA-2024-D-4488's robustness assessment requirement. The accuracy/stability dissociation—STER rising [X]–[X]× above baseline while Δaccuracy < [X%]—is the primary finding, confirmed across all stressor conditions and both hardware architectures. STER is proposed as a candidate addition to pre-market validation protocols for edge-deployed SaMD, complementing TFS (timing-layer monitoring) in a layered defense-in-depth architecture consistent with IEC 62304 and ISO 14971.

# References

[1] FDA, "Artificial Intelligence-Enabled Device Software Functions: Lifecycle Management and Marketing Submission Recommendations," Draft Guidance, Docket FDA-2024-D-4488, January 2025.

[2] FDA, "Good Machine Learning Practice for Medical Device Development: Guiding Principles," October 2021.

[3] [ADD: ACM TOIT multi-tenancy throughput paper — Reaching for the Sky, 2022]

[4] [ADD: GPU isolation / performance isolation for inference processes — arXiv:2601.07600, 2026]

[5] [ADD: DeepEdgeBench — Baller et al., ResearchGate, 2021]

[6] [ADD: Edge Devices Inference Performance Comparison — arXiv:2306.12093, 2023]

[7] [ADD: Increasing Safety of Neural Networks in Medical Devices — Springer SAFECOMP 2019]

[8] IEC 62304:2006+AMD1:2015, "Medical device software — Software life cycle processes," IEC, Geneva, 2015.

[9] ISO 14971:2019, "Medical devices — Application of risk management to medical devices," ISO, Geneva, 2019.

[10] [ADD: stress-ng tool citation — Paul E. McKenney]

[11] [ADD: Holoscan SDK / NVIDIA medical AI latency reference]

[12] [ADD: MobileNetV2 — Sandler et al., CVPR 2018]

[13] [ADD: TensorRT reference — NVIDIA Developer Documentation]

[14] [ADD: TFLite Edge TPU compiler reference — Google Coral Documentation]

# Appendix: Symbol Glossary

TABLE X. Symbol Definitions

| Symbol | Definition |
| --- | --- |
| T* | Output stability tolerance band (±5% relative deviation) — Eq. (2) |
| δᵢ | Per-inference output deviation — Eq. (1) |
| ȳ | Mean nominal output (E0 baseline) |
| STER | Safety-Threshold Exceedance Rate — Eq. (3) |
| Acc | Classification accuracy (top-1, N = 500 activations) |
| ε_acc | Accuracy detection sensitivity threshold (2%) |
| N | Number of inference activations per trial (500) |
| T_n | Nominal inference activation period (100 ms, 10 Hz) |
| f | Deployed inference model function |
| x | Fixed test input (held constant across all trials) |
| α | Statistical significance level (0.05) for z-test comparisons |

END OF DRAFT — All [X] placeholders to be replaced with experimental data before May 20, 2026 submission